## AGW (Bombas)
Code | Level | Class | Text | Short Text | Tractament 1 | Tractament 2 | Tractament 3
120 | MediumPriorityAlarm | Technical | Sequence number incorrect (lower) | Sequence number incorrect (lower) | 0 | 0 | 0
121 | MediumPriorityAlarm | Technical | Anomaly on AGW. Infusion parameters are not changed. Restart AGW. | Anomaly on AGW. Infusion parameters are not changed. Restart AGW. | 0 | 0 | 0
122 | MediumPriorityAlarm | Technical | AGW Version not supported: {0} {1} | AGW Version not supported: {0} {1} | 0 | 0 | 0
123 | MediumPriorityAlarm | Technical | Battery level low. {0} left | Battery level low. {0} left | 0 | 0 | 0
200 | ClinicalEvent | Physiological | Bolus | Bolus | 0 | 0 | 0
255 | UserEvent | Technical | Titration Event | Titration Event | 0 | 0 | 0
256 | UserEvent | Technical | Pump disconnected {0} {1} | Pump disconnected {0} {1} | 1 | 1 | 0
257 | UserEvent | Technical | Pump connected {0} {1} | Pump connected {0} {1} | 1 | 1 | 0
260 | UserEvent | Technical | Barcode read: {0} | Barcode read: {0} | 0 | 0 | 0
ENT 01 | HighPriorityAlarm | Technical | AC DISCONNECT | AC POWER FAIL | 0 | 0 | 0
ENT 02 | HighPriorityAlarm | Technical | DRIVE DISENGAGED | DRIVE DISENGAGED | 0 | 0 | 0
ENT 03 | HighPriorityAlarm | Technical | BATTERY EMPTY | BATTERY EMPTY | 1 | 1 | 0
ENT 04 | HighPriorityAlarm | Physiological | END OF INFUSION - KVO | END OF DELIVERY | 1 | 1 | 0
ENT 05 | HighPriorityAlarm | Physiological | END OF INFUSION - STOP | END OF DELIVERY | 1 | 1 | 0
ENT 06 | HighPriorityAlarm | Physiological | END OF INFUSION - KVO STOP | END OF DELIVERY | 1 | 1 | 0
ENT 07 | HighPriorityAlarm | Technical | BATTERY LOW | BATTERY LOW | 1 | 1 | 0
ENT 08 | HighPriorityAlarm | Physiological | NEAR END OF INFUSION | NEAR END OF DELIVERY | 1 | 0 | 0
ENT 09 | HighPriorityAlarm | Physiological | OCCLUSION | OCCLUSION | 1 | 1 | 0
ENT 10 | HighPriorityAlarm | Physiological | SYRINGE CLAMP OPEN | CHECK SYRINGE | 1 | 1 | 0
ENT 11 | HighPriorityAlarm | Physiological | TITRATION NOT CONFIRMED | TITRATION NOT CONFIRMED | 0 | 0 | 0
ENT 12 | HighPriorityAlarm | Physiological | VTBI DONE - CONTINUE | VTBI DONE | 0 | 0 | 0
ENT 13 | HighPriorityAlarm | Physiological | VTBI DONE - KVO | VTBI DONE | 0 | 0 | 0
ENT 14 | HighPriorityAlarm | Physiological | VTBI DONE - STOP | VTBI DONE - STOP | 0 | 0 | 0
ENT 15 | HighPriorityAlarm | Technical | ATTENTION | Attention (with " 3 beeps") | 0 | 0 | 0
ENT 17 | HighPriorityAlarm | Technical | COMMUNICATIONS TIMEOUT | COMMUNICATIONS TIMEOUT | 0 | 0 | 0
ENT 18 | HighPriorityAlarm | Physiological | INFUSION LINE DISCONNECTION | INFUSION LINE DISCONNECTION | 1 | 1 | 0
ENT 19 | HighPriorityAlarm | Physiological | EXCESSIVE RATE | EXCESSIVE RATE | 0 | 0 | 0
GW 01 | HighPriorityAlarm | Technical | BATTERY EMPTY | bAt (INTERNAL BATTERY DEPLETED) | 0 | 0 | 0
GW 02 | HighPriorityAlarm | Physiological | UPSTREAM OCCLUSION | Air OCCL (AIR-IN-LINE - UPSTREAM OCCLUSION) | 0 | 0 | 0
GW 03 | HighPriorityAlarm | Physiological | OCCLUSION | HI PrES (DOWNSTREAM OCCLUSION. A blockage has occurred downstream) | 0 | 0 | 0
GW 04 | HighPriorityAlarm | Physiological | DOOR OPEN | door (DOOR OPEN. Door was opened during an infusion) | 0 | 0 | 0
GW 05 | HighPriorityAlarm | Technical | WRONG SET | bAd SEt (INCORRECT INFUSION SET. Set incorrectly loaded or set worn) | 0 | 0 | 0
GW 06 | HighPriorityAlarm | Physiological | END OF INFUSION - KVO | End (Finished pre-set volume to be infused.) | 0 | 0 | 0
GW 07 | HighPriorityAlarm | Physiological | FLOW SENSOR DISCONNECTED | FLo SEnS (FLOW SENSOR CONNECTION ERROR) | 0 | 0 | 0
GW 08 | HighPriorityAlarm | Physiological | FLOW ERROR | FLo Err (FLOW ERROR) | 0 | 0 | 0
LVP 01 | MediumPriorityAlarm | Technical | AC DISCONNECT | AC POWER FAIL | 0 | 0 | 0
LVP 02 | HighPriorityAlarm | Technical | BATTERY EMPTY | BATTERY EMPTY | 1 | 1 | 0
LVP 03 | MediumPriorityAlarm | Technical | BATTERY LOW | BATTERY LOW | 1 | 1 | 0
LVP 04 | MediumPriorityAlarm | Physiological | TITRATION NOT CONFIRMED | TITRATION NOT CONFIRMED | 0 | 0 | 0
LVP 05 | MediumPriorityAlarm | Physiological | VTBI DONE - KVO | VTBI DONE | 0 | 0 | 0
LVP 06 | MediumPriorityAlarm | Technical | ATTENTION | Attention (with " 3 beeps") | 1 | 0 | 0
LVP 07 | HighPriorityAlarm | Physiological | OCCLUSION | DOWNSTREAM OCCLUSION | 0 | 0 | 0
LVP 08 | HighPriorityAlarm | Physiological | AIR IN LINE | AIR IN LINE | 1 | 1 | 0
LVP 09 | HighPriorityAlarm | Physiological | AIR IN LINE NOW | AIR-IN-LINE | 1 | 1 | 0
LVP 10 | HighPriorityAlarm | Technical | DOOR OPEN | DOOR OPEN | 0 | 0 | 0
LVP 11 | HighPriorityAlarm | Physiological | FLOW ERROR | FLOW ERROR | 0 | 0 | 0
LVP 12 | HighPriorityAlarm | Physiological | FREE FLOW | FREE FLOW | 0 | 0 | 0
LVP 13 | HighPriorityAlarm | Technical | FLOW SENSOR DISCONNECTED | FLOW SENSOR DISCONNECT | 0 | 0 | 0
LVP 14 | MediumPriorityAlarm | Physiological | SET VTBI | SET VTBI | 0 | 0 | 0
LVP 15 | MediumPriorityAlarm | Physiological | SET MISLOAD | SET MISLOAD | 0 | 0 | 0
LVP 16 | HighPriorityAlarm | Technical | SAFETY CLAMP | SAFETY CLAMP | 0 | 0 | 0
LVP 17 | HighPriorityAlarm | Technical | DOOR CLOSE INCOMPLETE | DOOR CLOSE INCOMPLETE | 0 | 0 | 0
LVP 18 | MediumPriorityAlarm | Technical | SET NOT FITTED | SET NOT FITTED | 0 | 0 | 0
LVP 19 | HighPriorityAlarm | Physiological | UPSTREAM OCCLUSION | UPSTREAM OCCLUSION | 1 | 1 | 0
LVP 20 | HighPriorityAlarm | Technical | WRONG SET | WRONG SET | 0 | 0 | 0
LVP 21 | HighPriorityAlarm | Technical | ACCUMULATED AIR IN LINE | AIR IN LINE | 1 | 1 | 0
LVP 22 | MediumPriorityAlarm | Physiological | KVO ATTENTION | KVO Attention | 0 | 0 | 0
LVP 23 | HighPriorityAlarm | Technical | LEVER OPEN | LEVER OPEN | 0 | 0 | 0
LVP 24 | HighPriorityAlarm | Technical | NO FLOW | NO FLOW | 0 | 0 | 0
LVP 25 | HighPriorityAlarm | Physiological | NO PRIMARY VTBI | NO PRIMARY VTBI | 0 | 0 | 0
LVP 26 | HighPriorityAlarm | Technical | PRIMARY FLOW DURING SECONDARY | PRIMARY FLOW DURING SECONDARY | 0 | 0 | 0
LVP 27 | MediumPriorityAlarm | Technical | SET CLOCK | SET CLOCK | 0 | 0 | 0
LVP 28 | MediumPriorityAlarm | Physiological | NEAR END OF INFUSION | NEAR END OF INFUSION | 1 | 0 | 0
SE 01 | HighPriorityAlarm | Physiological | ACCUMULATED AIR IN LINE | ACCUMULATED AIR IN LINE | 0 | 0 | 0
SE 02 | HighPriorityAlarm | Physiological | AIR IN LINE | AIR IN LINE | 1 | 1 | 0
SE 03 | MediumPriorityAlarm | Technical | ATTENTION | HOLD TIME EXCEEDED | 0 | 0 | 0
SE 04 | MediumPriorityAlarm | Technical | BATTERY LOW | Battery Low | 1 | 1 | 0
SE 05 | MediumPriorityAlarm | Technical | COMMUNICATION ALARM | COMPUTER LINK FAILURE | 0 | 0 | 0
SE 06 | HighPriorityAlarm | Physiological | FLOW ERROR | Resistance Alert | 0 | 0 | 0
SE 07 | HighPriorityAlarm | Technical | FLOW SENSOR DISCONNECTED | FLOW SENSOR UNPLUGGED | 0 | 0 | 0
SE 08 | HighPriorityAlarm | Physiological | LEVER OPEN | LATCH OPEN | 0 | 0 | 0
SE 09 | HighPriorityAlarm | Physiological | OCCLUSION | OCCLUSION DOWNSTREAM | 1 | 1 | 0
SE 10 | HighPriorityAlarm | Physiological | PRIMARY FLOW DURING SECONDARY | PRIMARY FLOW DETECTED DURING SECONDARY | 0 | 0 | 0
SE 11 | HighPriorityAlarm | Physiological | SET EMPTY | NO UPSTREAM FLOW DETECTED | 0 | 0 | 0
SE 12 | MediumPriorityAlarm | Physiological | SET NOT FITTED | SET OUT | 0 | 0 | 0
SE 13 | MediumPriorityAlarm | Technical | SETUP TIME EXCEEDED | SETUP TIME EXCEEDED | 1 | 1 | 0
SE 14 | HighPriorityAlarm | Technical | STUCK SWITCH | KEY STUCK | 0 | 0 | 0
SE 15 | HighPriorityAlarm | Physiological | UPSTREAM OCCLUSION | OCCLUSION UPSTREAM | 1 | 1 | 0
SE 16 | MediumPriorityAlarm | Physiological | VTBI DONE | VTBI = 0 | 0 | 0 | 0
SE 17 | MediumPriorityAlarm | Technical | COMPUTER CONTROL | COMPUTER CONTROL RELEASED | 0 | 0 | 0
SYR 01 | MediumPriorityAlarm | Technical | AC DISCONNECT | AC POWER FAIL | 0 | 0 | 0
SYR 02 | HighPriorityAlarm | Technical | DRIVE DISENGAGED | DRIVE DISENGAGED | 0 | 0 | 0
SYR 03 | HighPriorityAlarm | Technical | BATTERY EMPTY | BATTERY EMPTY | 1 | 1 | 0
SYR 04 | MediumPriorityAlarm | Physiological | END OF INFUSION - KVO | END OF INFUSION | 1 | 1 | 0
SYR 05 | MediumPriorityAlarm | Physiological | END OF INFUSION - STOP | END OF INFUSION | 1 | 1 | 0
SYR 06 | MediumPriorityAlarm | Physiological | END OF INFUSION - KVO STOP | END OF INFUSION | 1 | 1 | 0
SYR 07 | MediumPriorityAlarm | Technical | BATTERY LOW | BATTERY LOW | 1 | 1 | 0
SYR 08 | MediumPriorityAlarm | Physiological | NEAR END OF INFUSION | NEAR END OF INFUSION | 1 | 0 | 0
SYR 09 | HighPriorityAlarm | Physiological | OCCLUSION | OCCLUSION | 1 | 1 | 0
SYR 10 | HighPriorityAlarm | Technical | SYRINGE CLAMP OPEN | CHECK SYRINGE | 1 | 1 | 0
SYR 11 | MediumPriorityAlarm | Technical | TITRATION NOT CONFIRMED | TITRATION NOT CONFIRMED | 0 | 0 | 0
SYR 12 | HighPriorityAlarm | Physiological | VTBI DONE - CONTINUE | VTBI DONE | 0 | 0 | 0
SYR 13 | HighPriorityAlarm | Physiological | VTBI DONE - KVO | VTBI DONE | 0 | 0 | 0
SYR 14 | HighPriorityAlarm | Physiological | VTBI DONE - STOP | VTBI DONE - STOP | 0 | 0 | 0
SYR 15 | MediumPriorityAlarm | Technical | ATTENTION | Attention (with " 3 beeps") | 1 | 0 | 0
SYR 16 | HighPriorityAlarm | Technical | PRESSURE DISC REMOVED | PRESSURE DISC OUT | 0 | 0 | 0
SYR 17 | MediumPriorityAlarm | Technical | COMMUNICATIONS TIMEOUT | COMMUNICATIONS TIMEOUT | 0 | 0 | 0
SYR 18 | HighPriorityAlarm | Technical | INFUSION LINE DISCONNECTION | INFUSION LINE DISCONNECTION | 1 | 1 | 0
SYR 19 | HighPriorityAlarm | Physiological | EXCESSIVE RATE | EXCESSIVE RATE | 0 | 0 | 0
SYR 20 | HighPriorityAlarm | Technical | TCI TAG NOT RECOGNISED | TCI TAG NOT RECOGNISED | 0 | 0 | 0